{"text_format":"1","workbook":"xlsx","n_sheets":2,"sheets":["Info & Summary","Current Calibrations"],"not_rendered":{"charts":2,"images":9,"macros":0,"external_links":0,"defined_names":1,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"Info & Summary","index":0,"n_rows":17,"n_cols":14,"rows_kept":17,"cut":false,"hidden":false,"merged":[[14,0,14,7],[15,0,15,7],[7,0,7,1],[8,0,8,1],[9,0,9,1],[12,0,12,1],[13,0,13,1],[10,0,10,13],[11,2,11,3],[11,4,11,5],[11,6,11,7],[11,8,11,9],[11,10,11,11],[11,12,11,13],[4,0,4,7],[5,0,5,7],[6,0,6,1],[6,2,6,3],[6,4,6,5],[6,6,6,7],[0,0,0,7],[1,1,1,2],[1,3,1,4],[1,5,1,6],[2,1,2,2],[2,3,2,4],[2,5,2,6]]}
["Instrument Information"]
[null,"Instrument Type",null,"Serial Number",null,"Instrument Name"]
[null,"SBP",null,"enter S/N",null,"enter name"]
[]
["Current Calibration Summary"]
["Pressure Cells"]
[null,null,"TPC",null,"PPC A",null,"PPC B"]
["Zero",null,{"$f":"'Current Calibrations'!$B$24","v":-1398.3916943521594},{"$f":"'Current Calibrations'!$C$24","v":"mV"},{"$f":"'Current Calibrations'!$E$24","v":-30.252159468438435},{"$f":"'Current Calibrations'!$F$24","v":"mV"},{"$f":"'Current Calibrations'!$H$24","v":-1878.9189368770762},{"$f":"'Current Calibrations'!$I$24","v":"mV"}]
["Slope",null,{"$f":"'Current Calibrations'!$B$25","v":2.776664451827243},{"$f":"'Current Calibrations'!$C$25","v":"mV/psi"},{"$f":"'Current Calibrations'!$E$25","v":1.6178272425249169},{"$f":"'Current Calibrations'!$F$25","v":"mV/psi"},{"$f":"'Current Calibrations'!$H$25","v":1.6854850498338874},{"$f":"'Current Calibrations'!$I$25","v":"mV/psi"}]
["Sensitivity",null,{"$f":"'Current Calibrations'!$B$26","v":402.72097242358586},{"$f":"'Current Calibrations'!$C$26","v":"mV/MPa"},{"$f":"'Current Calibrations'!$E$26","v":234.64591117383594},{"$f":"'Current Calibrations'!$F$26","v":"mV/MPa"},{"$f":"'Current Calibrations'!$H$26","v":244.45884263322978},{"$f":"'Current Calibrations'!$I$26","v":"mV/MPa"}]
["Arm Springs"]
[null,null,"Arm 1",null,"Arm 2",null,"Arm 3",null,"Arm 4",null,"Arm 5",null,"Arm 6"]
["Zero",null,{"$f":"'Current Calibrations'!$B$45","v":-280.65614617940196},{"$f":"'Current Calibrations'!$C$45","v":"mV"},{"$f":"'Current Calibrations'!$E$45","v":246.77275747508293},{"$f":"'Current Calibrations'!$F$45","v":"mV"},{"$f":"'Current Calibrations'!$H$45","v":-266.86976744186063},{"$f":"'Current Calibrations'!$I$45","v":"mV"},{"$f":"'Current Calibrations'!$K$45","v":-280.65614617940196},{"$f":"'Current Calibrations'!$C$45","v":"mV"},{"$f":"'Current Calibrations'!$N$45","v":246.77275747508293},{"$f":"'Current Calibrations'!$F$45","v":"mV"},{"$f":"'Current Calibrations'!$Q$45","v":-266.86976744186063},{"$f":"'Current Calibrations'!$I$45","v":"mV"}]
["Slope",null,{"$f":"'Current Calibrations'!$B$46","v":302.65083056478403},{"$f":"'Current Calibrations'!$C$46","v":"mV/mm"},{"$f":"'Current Calibrations'!$E$46","v":288.78205980066446},{"$f":"'Current Calibrations'!$F$46","v":"mV/mm"},{"$f":"'Current Calibrations'!$H$46","v":291.2418604651163},{"$f":"'Current Calibrations'!$I$46","v":"mV/mm"},{"$f":"'Current Calibrations'!$K$46","v":302.65083056478403},{"$f":"'Current Calibrations'!$C$46","v":"mV/mm"},{"$f":"'Current Calibrations'!$N$46","v":288.78205980066446},{"$f":"'Current Calibrations'!$F$46","v":"mV/mm"},{"$f":"'Current Calibrations'!$Q$46","v":291.2418604651163},{"$f":"'Current Calibrations'!$I$46","v":"mV/mm"}]
["Straightness Data Summary"]
["See the straightness calibration sheet for the locations of points A to F"]
["Date","A","B","C","D","E","F","Operator"]

{"sheet":"Current Calibrations","index":1,"n_rows":108,"n_cols":19,"rows_kept":108,"cut":false,"hidden":false,"merged":[[1,2,1,3],[1,4,1,5],[1,6,1,7],[1,8,1,9],[28,0,28,18],[0,0,0,1],[0,2,0,3],[0,4,0,5],[0,6,0,7],[2,0,5,0],[2,1,5,9],[6,4,6,5],[6,6,6,7],[6,8,6,9],[27,0,27,18],[0,8,0,9],[1,0,1,1]]}
["Calibration Date",null,"Operator",null,"Instrument Type",null,"Serial Number",null,"Instrument Name"]
[null,null,null,null,{"$f":"'Info & Summary'!B3","v":"SBP"},null,{"$f":"'Info & Summary'!D3","v":"enter S/N"},null,{"$f":"'Info & Summary'!F3","v":"enter name"}]
["Notes:"]
[]
[]
[]
[null,null,null,null,"Pressure cells",null,"Test Gauge:",null,"10759918 (0-600psi)"]
[]
["psi","TPC","Linearity","Hysteresis","PPC A","Linearity","Hysteresis","PPC B","Linearity","Hysteresis"]
[null,"(mV)","(%)","(%)","(mV)","(%)","(%)","(mV)","(%)","(%)"]
[0,-1400,{"$f":"(B12-B11)/B$25","v":98.391435025652811},{"$f":"(B11-B23)*100/(B$17-B$11)","v":-0.34223956769739089},-30.1,{"$f":"(E12-E11)/E$25","v":98.898075019611227},{"$f":"(E11-E23)*100/(E$17-E$11)","v":-9.2908021059151649E-2},-1883.1,{"$f":"(H12-H11)/H$25","v":102.10710561743706},{"$f":"(H11-H23)*100/(H$17-H$11)","v":-0.50380322038920378}]
[100,-1126.8,99.868026839727364,{"$f":"(B12-B22)*100/(B$17-B$11)","v":-0.55238667066947134},129.9,100.01067836358185,{"$f":"(E12-E22)*100/(E$17-E$11)","v":-0.14452358831423615},-1711,100.1492122499906,{"$f":"(H12-H22)*100/(H$17-H$11)","v":-0.17781290131383529}]
[200,-849.5,100.8044021364576,{"$f":"(B13-B21)*100/(B$17-B$11)","v":-0.63644551185830223},291.7,100.99965911377798,{"$f":"(E13-E21)*100/(E$17-E$11)","v":-0.18581604211830405},-1542.2,100.80184337247285,{"$f":"(H13-H21)*100/(H$17-H$11)","v":-0.17781290131383529}]
[300,-569.6,100.37222892258212,{"$f":"(B14-B20)*100/(B$17-B$11)","v":-0.55839087361153217},455.1,100.38154614490539,{"$f":"(E14-E20)*100/(E$17-E$11)","v":-9.2908021059149082E-2},-1372.3,99.852561739771431,{"$f":"(H14-H20)*100/(H$17-H$11)","v":-7.9027956139479855E-2}]
[400,-290.89999999999998,100.30020005360286,{"$f":"(B15-B19)*100/(B$17-B$11)","v":-0.52236565595917073},617.5,99.577999285371064,{"$f":"(E15-E19)*100/(E$17-E$11)","v":-8.258490760812992E-2},-1204,99.199930617289311,{"$f":"(H15-H19)*100/(H$17-H$11)","v":-8.890645065692887E-2}]
[500,-12.4,100.08411344666511,{"$f":"(B16-B18)*100/(B$17-B$11)","v":-0.37826478534974489},778.6,98.898075019611227,{"$f":"(E16-E18)*100/(E$17-E$11)","v":-0.11355424796118743},-1036.8,98.48796939276329,{"$f":"(H16-H18)*100/(H$17-H$11)","v":-9.8784945174355432E-2}]
[600,265.5,-97.815204073818819,null,938.6,-98.21815075385139,null,-870.8,-97.894668372324958]
[500,-6.1,-99.435853625851877,null,779.7,-99.763433176032876,null,-1035.8,-99.259260719333085]
[400,-282.2,-100.15614231564436,null,618.29999999999995,-100.31973484801811,null,-1203.0999999999999,-99.911891841815333]
[300,-560.29999999999995,-100.33621448809249,null,456,-100.44335744179266,null,-1371.5,-100.2085423520345]
[200,-838.9,-100.37222892258208,null,293.5,-100.25792355113087,null,-1540.4,-100.1492122499906]
[100,-1117.5999999999999,-99.651940232789642,null,131.30000000000001,-99.20713150404751,null,-1709.2,-100.1492122499906]
[0,-1394.3,null,null,-29.2,null,null,-1878]
["Zero",{"$f":"INTERCEPT(B11:B23,$A11:$A23)","v":-1398.3916943521594},"mV",null,{"$f":"INTERCEPT(E11:E23,$A11:$A23)","v":-30.252159468438435},"mV",null,{"$f":"INTERCEPT(H11:H23,$A11:$A23)","v":-1878.9189368770762},"mV"]
["Slope",{"$f":"SLOPE(B11:B23,$A11:$A23)","v":2.776664451827243},"mV/psi",null,{"$f":"SLOPE(E11:E23,$A11:$A23)","v":1.6178272425249169},"mV/psi",null,{"$f":"SLOPE(H11:H23,$A11:$A23)","v":1.6854850498338874},"mV/psi"]
["Sensitivity",{"$f":"B25*1000/6.89476","v":402.72097242358586},"mV/MPa",null,{"$f":"E25*1000/6.89476","v":234.64591117383594},"mV/MPa",null,{"$f":"H25*1000/6.89476","v":244.45884263322978},"mV/MPa"]
[]
["Arm Springs"]
[]
["mm","Arm 1","Linearity","Hysteresis","Arm 2","Linearity","Hysteresis","Arm 3","Linearity","Hysteresis","Arm 4","Linearity","Hysteresis","Arm 5","Linearity","Hysteresis","Arm 6","Linearity","Hysteresis"]
[null,"(mV)","(%)","(%)","(mV)","(%)","(%)","(mV)","(%)","(%)","(mV)","(%)","(%)","(mV)","(%)","(%)","(mV)","(%)","(%)"]
[0,-290.5,{"$f":"(B33-B32)*100/B$46","v":104.04729417472851},{"$f":"(B32-B44)*100/(B$38-B$32)","v":0},244.5,{"$f":"(E33-E32)*100/E$46","v":101.84150642979913},{"$f":"(E32-E44)*100/(E$38-E$32)","v":0.2487850034714193},-274.8,{"$f":"(H33-H32)*100/H$46","v":104.14024945302393},{"$f":"(H32-H44)*100/(H$38-H$32)","v":0},-290.5,{"$f":"(K33-K32)*100/K$46","v":104.04729417472851},{"$f":"(K32-K44)*100/(K$38-K$32)","v":0},244.5,{"$f":"(N33-N32)*100/N$46","v":101.84150642979913},{"$f":"(N32-N44)*100/(N$38-N$32)","v":0.2487850034714193},-274.8,{"$f":"(Q33-Q32)*100/Q$46","v":104.14024945302393},{"$f":"(Q32-Q44)*100/(Q$38-Q$32)","v":0}]
[1,24.4,102.72563912011147,{"$f":"(B33-B43)*100/(B$38-B$32)","v":0.11003521126760564},538.6,100.83728892335088,{"$f":"(E33-E43)*100/(E$38-E$32)","v":0.20249942143022448},28.5,101.42772729450466,{"$f":"(H33-H43)*100/(H$38-H$32)","v":0.26920213070622601},24.4,102.72563912011147,{"$f":"(K33-K43)*100/(K$38-K$32)","v":0.11003521126760564},538.6,100.83728892335088,{"$f":"(N33-N43)*100/(N$38-N$32)","v":0.20249942143022448},28.5,101.42772729450466,{"$f":"(Q33-Q43)*100/(Q$38-Q$32)","v":0.26920213070622601}]
[2,335.3,98.99196359081823,{"$f":"(B34-B42)*100/(B$38-B$32)","v":0.20906690140845133},829.8,99.902327796657744,{"$f":"(E34-E42)*100/(E$38-E$32)","v":0.37607035408470257},323.89999999999998,100.05429835348227,{"$f":"(H34-H42)*100/(H$38-H$32)","v":0.3952116386963731},335.3,98.99196359081823,{"$f":"(K34-K42)*100/(K$38-K$32)","v":0.20906690140845133},829.8,99.902327796657744,{"$f":"(N34-N42)*100/(N$38-N$32)","v":0.37607035408470257},323.89999999999998,100.05429835348227,{"$f":"(Q34-Q42)*100/(Q$38-Q$32)","v":0.3952116386963731}]
[3,634.9,98.529384321702281,{"$f":"(B35-B41)*100/(B$38-B$32)","v":0.15955105633802694},1118.3,99.659930467515068,{"$f":"(E35-E41)*100/(E$38-E$32)","v":0.28928488775746353},615.29999999999995,98.818212306562131,{"$f":"(H35-H41)*100/(H$38-H$32)","v":0.37802852397043984},634.9,98.529384321702281,{"$f":"(K35-K41)*100/(K$38-K$32)","v":0.15955105633802694},1118.3,99.659930467515068,{"$f":"(N35-N41)*100/(N$38-N$32)","v":0.28928488775746353},615.29999999999995,98.818212306562131,{"$f":"(Q35-Q41)*100/(Q$38-Q$32)","v":0.37802852397043984}]
[4,933.1,98.165929181682586,{"$f":"(B36-B40)*100/(B$38-B$32)","v":0.24207746478873116},1406.1,99.001994859842142,{"$f":"(E36-E40)*100/(E$38-E$32)","v":0.29507058551260756},903.1,97.89114777137199,{"$f":"(H36-H40)*100/(H$38-H$32)","v":0.2291081963457243},933.1,98.165929181682586,{"$f":"(K36-K40)*100/(K$38-K$32)","v":0.24207746478873116},1406.1,99.001994859842142,{"$f":"(N36-N40)*100/(N$38-N$32)","v":0.29507058551260756},903.1,97.89114777137199,{"$f":"(Q36-Q40)*100/(Q$38-Q$32)","v":0.2291081963457243}]
[5,1230.2,98.099846428951665,{"$f":"(B37-B39)*100/(B$38-B$32)","v":0.10453345070423035},1692,97.270585365965914,{"$f":"(E37-E39)*100/(E$38-E$32)","v":0.16778523489933411},1188.2,97.135761853809612,{"$f":"(H37-H39)*100/(H$38-H$32)","v":0.10882639326422425},1230.2,98.099846428951665,{"$f":"(K37-K39)*100/(K$38-K$32)","v":0.10453345070423035},1692,97.270585365965914,{"$f":"(N37-N39)*100/(N$38-N$32)","v":0.16778523489933411},1188.2,97.135761853809612,{"$f":"(Q37-Q39)*100/(Q$38-Q$32)","v":0.10882639326422425}]
[6,1527.1,-98.727632579894816,null,1972.9,-98.27480287241417,null,1471.1,-97.788140600795288,null,1527.1,-98.727632579894816,null,1972.9,-98.27480287241417,null,1471.1,-97.788140600795288]
[5,1228.3,-98.991963590818216,null,1689.1,-99.763815037147623,null,1186.3,-98.612197965408726,null,1228.3,-98.991963590818216,null,1689.1,-99.763815037147623,null,1186.3,-98.612197965408726]
[4,928.7,-98.033763676220872,null,1401,-99.625302277637559,null,899.1,-99.710941118226657,null,928.7,-98.033763676220872,null,1401,-99.625302277637559,null,899.1,-99.710941118226657]
[3,632,-99.289335978107076,null,1113.3,-100.42175064482061,null,608.70000000000005,-100.15730552405897,null,632,-99.289335978107076,null,1113.3,-100.42175064482061,null,608.70000000000005,-100.15730552405897]
[2,331.5,-102.13089434553378,null,823.3,-99.798443227025146,null,317,-100.67234137694236,null,331.5,-102.13089434553378,null,823.3,-99.798443227025146,null,317,-100.67234137694236]
[1,22.4,-103.38646664741997,null,535.1,-102.11853194881931,null,23.8,-102.52647044732261,null,22.4,-103.38646664741997,null,535.1,-102.11853194881931,null,23.8,-102.52647044732261]
[0,-290.5,null,null,240.2,null,null,-274.8,null,null,-290.5,null,null,240.2,null,null,-274.8]
["Zero",{"$f":"INTERCEPT(B32:B44,$A32:$A44)","v":-280.65614617940196},"mV",null,{"$f":"INTERCEPT(E32:E44,$A32:$A44)","v":246.77275747508293},"mV",null,{"$f":"INTERCEPT(H32:H44,$A32:$A44)","v":-266.86976744186063},"mV",null,{"$f":"INTERCEPT(K32:K44,$A32:$A44)","v":-280.65614617940196},"mV",null,{"$f":"INTERCEPT(N32:N44,$A32:$A44)","v":246.77275747508293},"mV",null,{"$f":"INTERCEPT(Q32:Q44,$A32:$A44)","v":-266.86976744186063},"mV"]
["Slope",{"$f":"SLOPE(B32:B44,$A32:$A44)","v":302.65083056478403},"mV/mm",null,{"$f":"SLOPE(E32:E44,$A32:$A44)","v":288.78205980066446},"mV/mm",null,{"$f":"SLOPE(H32:H44,$A32:$A44)","v":291.2418604651163},"mV/mm",null,{"$f":"SLOPE(K32:K44,$A32:$A44)","v":302.65083056478403},"mV/mm",null,{"$f":"SLOPE(N32:N44,$A32:$A44)","v":288.78205980066446},"mV/mm",null,{"$f":"SLOPE(Q32:Q44,$A32:$A44)","v":291.2418604651163},"mV/mm"]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
["Do Not Enter Text Below This Line.",null,null,null,"Do Not Enter Text Below This Line.",null,null,null,"Do Not Enter Text Below This Line.",null,null,null,"Do Not Enter Text Below This Line.",null,null,null,"Do Not Enter Text Below This Line."]
[]
[]
[]
["mm","Arm 1","Arm 2","Arm 3","Arm 4","Arm 5","Arm 6",null,"psi","TPC","PPC A","PPC B"]
["Plot?",true,true,true,true,true,true,null,"plot?",true,true,true]
[0,{"$f":"IF(B$85=TRUE,B32,#N/A)","v":-290.5},{"$f":"IF(C$85=TRUE,E32,#N/A)","v":244.5},{"$f":"IF(D$85=TRUE,H32,#N/A)","v":-274.8},{"$f":"IF(E$85=TRUE,K32,#N/A)","v":-290.5},{"$f":"IF(F$85=TRUE,N32,#N/A)","v":244.5},{"$f":"IF(G$85=TRUE,Q32,#N/A)","v":-274.8},null,0,{"$f":"IF(J$85=TRUE,B11,#N/A)","v":-1400},{"$f":"IF(K$85=TRUE,E11,#N/A)","v":-30.1},{"$f":"IF(L$85=TRUE,H11,#N/A)","v":-1883.1}]
[1,24.4,538.6,28.5,24.4,538.6,28.5,null,100,{"$f":"IF(J$85=TRUE,B12,#N/A)","v":-1126.8},{"$f":"IF(K$85=TRUE,E12,#N/A)","v":129.9},{"$f":"IF(L$85=TRUE,H12,#N/A)","v":-1711}]
[2,335.3,829.8,323.89999999999998,335.3,829.8,323.89999999999998,null,200,-849.5,291.7,-1542.2]
[3,634.9,1118.3,615.29999999999995,634.9,1118.3,615.29999999999995,null,300,-569.6,455.1,-1372.3]
[4,933.1,1406.1,903.1,933.1,1406.1,903.1,null,400,-290.89999999999998,617.5,-1204]
[5,1230.2,1692,1188.2,1230.2,1692,1188.2,null,500,-12.4,778.6,-1036.8]
[6,1527.1,1972.9,1471.1,1527.1,1972.9,1471.1,null,600,265.5,938.6,-870.8]
[5,1228.3,1689.1,1186.3,1228.3,1689.1,1186.3,null,500,-6.1,779.7,-1035.8]
[4,928.7,1401,899.1,928.7,1401,899.1,null,400,-282.2,618.29999999999995,-1203.0999999999999]
[3,632,1113.3,608.70000000000005,632,1113.3,608.70000000000005,null,300,-560.29999999999995,456,-1371.5]
[2,331.5,823.3,317,331.5,823.3,317,null,200,-838.9,293.5,-1540.4]
[1,22.4,535.1,23.8,22.4,535.1,23.8,null,100,-1117.5999999999999,131.30000000000001,-1709.2]
[0,-290.5,240.2,-274.8,-290.5,240.2,-274.8,null,0,-1394.3,-29.2,-1878]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[null,null,null,null,false]